## Slide 1
ISO
13485
SKILLMED MASTER PRESENTATION
ISO 13485:2016 Medical Devices Quality Management Systems
Requirements for regulatory purposes — clause-by-clause guide with audit insights
www.skillmed.in • Free training & certification
Clauses 4 – 8
Design • Production • CAPA

## Slide 2
Introduction
Purpose, Scope & Business Value
SkillMed | ISO 13485:2016 Master Presentation Foundations
2
▸ Specifies QMS requirements for organizations involved in any life-cycle stage of a medical device
▸ Applies to manufacturers, suppliers, distributors, importers, service providers and sterilizers
▸ Primary aim: consistently meet customer AND applicable regulatory requirements
▸ Emphasises safety and performance — not continual improvement for its own sake
▸ Recognised basis for CE (EU MDR), MDSAP, Health Canada, FDA QMSR and CDSCO
Business value
▸ Market access in 100+ countries
▸ Fewer recalls & complaints
▸ Faster regulator audits
▸ Customer & tender trust

## Slide 3
Comparison
ISO 13485 vs ISO 9001
SkillMed | ISO 13485:2016 Master Presentation Foundations
3
ISO 9001:2015
▸ Generic — any industry
▸ Customer satisfaction & continual improvement
▸ Annex SL high-level structure
▸ Documented information (flexible)
▸ Risk-based thinking (general)
ISO 13485:2016
▸ Medical devices only
▸ Regulatory compliance & effectiveness
▸ Clause structure from ISO 9001:2008
▸ Many mandatory procedures & records
▸ Risk management across product realization

## Slide 4
Regulations
Global Regulatory Mapping
SkillMed | ISO 13485:2016 Master Presentation Foundations
4
Market | How ISO 13485 is used | Key link
European Union | Harmonised as EN ISO 13485:2016+A11:2021 | Presumption of conformity with MDR/IVDR QMS
United States | Incorporated by reference in FDA QMSR | 21 CFR Part 820 (effective Feb 2026)
MDSAP countries | Audit model built on ISO 13485 | USA, Canada, Brazil, Japan, Australia
India | QMS aligned via MDR 2017 Fifth Schedule | CDSCO manufacturing licence
Canada | Required via MDSAP certificate | Class II–IV device licences

## Slide 5
Clauses 0–8
Structure of the Standard
SkillMed | ISO 13485:2016 Master Presentation Foundations
5
1
Clauses 0–3
Introduction, scope, normative references, terms & definitions
2
Clause 4
Quality management system & documentation
3
Clause 5
Management responsibility
4
Clause 6
Resource management
5
Clause 7
Product realization — the largest clause
6
Clause 8
Measurement, analysis & improvement

## Slide 6
Clause 0.2 / 4.1.2
Process Approach & Risk-Based Thinking
SkillMed | ISO 13485:2016 Master Presentation Foundations
6
▸ Manage activities as interrelated processes with inputs, outputs and controls
▸ Apply a risk-based approach to control of QMS processes
▸ Risk in ISO 13485 = risk to safety & performance of the device and to regulatory compliance
▸ Product risk management is performed per ISO 14971
▸ Effort and controls are proportionate to the risk
Remember
▸ 'Documented' = established, implemented & maintained
▸ 'Where appropriate' must be justified
▸ Regulatory requirements are always in scope

## Slide 7
01
Quality Management System
Clause 4 — general requirements, documentation, Medical Device File, control of documents and records
7

## Slide 8
4.1
General QMS Requirements
SkillMed | ISO 13485:2016 Master Presentation Clause 4
8
▸ Document the organization's role(s) under applicable regulations
▸ Determine processes and their sequence & interaction
▸ Ensure resources, monitoring, measurement and analysis
▸ Control outsourced processes proportionate to risk, via written quality agreements
▸ Validate computer software used in the QMS before first use and after changes
Outsourcing tips
▸ You stay responsible for outsourced work
▸ Define controls in agreements
▸ Monitor supplier performance

## Slide 9
4.2.1
Documentation Hierarchy
SkillMed | ISO 13485:2016 Master Presentation Clause 4
9
Quality Manual
Scope, exclusions, process interaction
Policies
Quality policy & objectives
Procedures
SOPs required by the standard and regulations
Work Instructions
Step-by-step task guidance
Records
Objective evidence of activities
Every level must be controlled — documents per 4.2.4, records per 4.2.5

## Slide 10
4.2.2
The Quality Manual
SkillMed | ISO 13485:2016 Master Presentation Clause 4
10
▸ Scope of the QMS, with details and justification of exclusions or non-applications
▸ Documented procedures or references to them
▸ Description of the interaction between QMS processes
▸ Outline of the documentation structure
▸ Exclusions limited to design & development only when regulations permit
Auditor checks
▸ Is the scope consistent with certificate?
▸ Are exclusions justified?
▸ Is the process map current?

## Slide 11
4.2.3
Medical Device File (MDF)
SkillMed | ISO 13485:2016 Master Presentation Clause 4
11
1
Description
General description, intended use/purpose and labelling incl. IFU
2
Specifications
Product specifications
3
Manufacturing
Specifications/procedures for manufacturing, packaging, storage, handling, distribution
4
Measuring
Procedures for measuring & monitoring
5
Installation
Requirements for installation, as appropriate
6
Servicing
Procedures for servicing, as appropriate

## Slide 12
4.2.4
Control of Documents
SkillMed | ISO 13485:2016 Master Presentation Clause 4
12
Review & approve
Before issue, by authorised persons
Update
Re-review and re-approve changes
Identify
Current revision status visible
Available
At points of use, legible
Obsolete
Prevent unintended use; retain
Retain obsolete documents for at least the lifetime of the device, or as regulations require

## Slide 13
4.2.5
Control of Records
SkillMed | ISO 13485:2016 Master Presentation Clause 4
13
▸ Records remain legible, readily identifiable and retrievable
▸ Documented procedure for identification, storage, security, integrity, retrieval, retention and disposition
▸ Protect confidential health information per applicable regulations
▸ Changes to records must remain identifiable
▸ Retain for device lifetime, but not less than two years from release
Typical records
▸ Training
▸ Batch / DHR
▸ Calibration
▸ Complaints & CAPA
▸ Audit reports

## Slide 14
4.1.6 / 4.2.5
Electronic Records & QMS Software
SkillMed | ISO 13485:2016 Master Presentation Clause 4
14
Validation expectations
▸ Validate QMS software before use
▸ Approach proportionate to risk
▸ Re-validate after changes
▸ Keep validation records
Data integrity (ALCOA+)
▸ Attributable & legible
▸ Contemporaneous & original
▸ Accurate & complete
▸ Secure audit trails (cf. 21 CFR Part 11)

## Slide 15
Summary
Clause 4 Checklist
SkillMed | ISO 13485:2016 Master Presentation Clause 4
15
Item | Evidence an auditor expects
QMS roles | Documented role: manufacturer, importer, distributor…
Process map | Interactions and outsourced processes
Quality manual | Scope, exclusions, references
MDF | One file per device type/family
Document control | SOP, master list, change history
Records control | Retention matrix, backups
Software validation | Protocols & reports for QMS tools

## Slide 16
02
Management Responsibility
Clause 5 — commitment, policy, objectives, responsibility and management review
16

## Slide 17
5.1 / 5.2
Commitment & Customer Focus
SkillMed | ISO 13485:2016 Master Presentation Clause 5
17
▸ Top management communicates the importance of customer and regulatory requirements
▸ Establishes quality policy and objectives
▸ Conducts management reviews and ensures resources
▸ Ensures customer and applicable regulatory requirements are determined and met
Evidence
▸ Signed policy
▸ Review minutes
▸ Budget approvals
▸ Communication records

## Slide 18
5.3 / 5.4
Quality Policy & Objectives
SkillMed | ISO 13485:2016 Master Presentation Clause 5
18
Quality policy
▸ Applicable to the purpose
▸ Commitment to comply & maintain effectiveness
▸ Framework for objectives
▸ Communicated and understood
▸ Reviewed for suitability
Quality objectives (SMART)
▸ Measurable & consistent with policy
▸ e.g. Complaint closure ≤ 30 days
▸ e.g. Supplier OTIF ≥ 95%
▸ e.g. CAPA effectiveness ≥ 90%
▸ Planned at relevant functions

## Slide 19
5.5
Responsibility, Authority & Communication
SkillMed | ISO 13485:2016 Master Presentation Clause 5
19
1
Defined roles
Responsibilities and authorities documented and communicated
2
Interrelation
Document interrelation of staff who manage, perform and verify work affecting quality
3
Management Rep
Member of management ensuring processes are documented and reported
4
Communication
Internal communication on QMS effectiveness

## Slide 20
5.6.2
Management Review Inputs
SkillMed | ISO 13485:2016 Master Presentation Clause 5
20
# | Input | # | Input
1 | Feedback | 7 | Monitoring & measurement of processes
2 | Complaint handling | 8 | Monitoring & measurement of product
3 | Reporting to regulators | 9 | Corrective action
4 | Audits | 10 | Preventive action
5 | Follow-up from previous reviews | 11 | Changes affecting the QMS
6 | Recommendations for improvement | 12 | New or revised regulatory requirements

## Slide 21
5.6.3
Management Review Outputs
SkillMed | ISO 13485:2016 Master Presentation Clause 5
21
▸ Improvement needed to maintain suitability, adequacy and effectiveness of the QMS
▸ Improvement of product related to customer requirements
▸ Changes needed to respond to new or revised regulatory requirements
▸ Resource needs
▸ Review at documented planned intervals; keep records
Good practice
▸ Data-driven KPI pack
▸ Action owners & due dates
▸ Trend over last periods

## Slide 22
Audit pitfalls
Common Clause 5 Findings
SkillMed | ISO 13485:2016 Master Presentation Clause 5
22
1
Missing inputs
Review minutes skip one or more of the 12 required inputs
2
No decisions
Outputs not recorded as decisions and actions
3
Vague objectives
Objectives not measurable or not tracked
4
Interval missed
Reviews not held at the planned frequency
5
Policy unknown
Staff cannot explain how policy relates to their work
6
Regulatory gaps
Regulatory changes not discussed

## Slide 23
03
Resource Management
Clause 6 — people, competence, infrastructure and work environment
23

## Slide 24
6.1 / 6.2
Resources & Human Competence
SkillMed | ISO 13485:2016 Master Presentation Clause 6
24
▸ Determine and provide resources to implement the QMS and meet regulatory requirements
▸ Personnel competent on education, training, skills and experience
▸ Document the process for establishing competence, providing training and ensuring awareness
▸ Evaluate effectiveness of actions taken — proportionate to risk of the work
Records
▸ Job descriptions
▸ Competence matrix
▸ Training records
▸ Effectiveness checks

## Slide 25
6.2
Training Program Lifecycle
SkillMed | ISO 13485:2016 Master Presentation Clause 6
25
Identify need
Role, SOP changes, gaps
Plan
Curriculum & schedule
Deliver
Classroom, OJT, e-learning
Evaluate
Test, observation, KPIs
Record
Retain evidence
Effectiveness is the key word — attendance alone is not proof of competence

## Slide 26
6.3
Infrastructure
SkillMed | ISO 13485:2016 Master Presentation Clause 6
26
1
Buildings & space
Adequate for product conformity, preventing mix-ups
2
Process equipment
Hardware and software with maintenance requirements
3
Supporting services
Transport, communication, IT systems
4
Maintenance
Documented maintenance activities, intervals and records

## Slide 27
6.4.1
Work Environment & Cleanrooms
SkillMed | ISO 13485:2016 Master Presentation Clause 6
27
▸ Document requirements for the work environment needed for conformity
▸ Document health, cleanliness and clothing requirements for personnel if contact could affect product
▸ Cleanroom classification and monitoring per ISO 14644 where applicable
▸ Train or supervise personnel working temporarily in special conditions
Monitor
▸ Particles
▸ Pressure differentials
▸ Temperature & humidity
▸ Microbial counts

## Slide 28
6.4.2
Contamination Control
SkillMed | ISO 13485:2016 Master Presentation Clause 6
28
Contaminated product
▸ Plan & document arrangements for contaminated or potentially contaminated product
▸ Prevent contamination of work environment, personnel or product
▸ Returned and used devices
Sterile devices
▸ Document requirements for microorganism / particulate control
▸ Maintain cleanliness during assembly and packaging
▸ Bioburden monitoring before sterilization

## Slide 29
04
Product Realization — Planning & Design
Clause 7.1 – 7.3 — planning, customer processes and design controls
29

## Slide 30
7.1
Planning of Product Realization
SkillMed | ISO 13485:2016 Master Presentation Clause 7
30
▸ Plan and develop processes needed for product realization
▸ Document one or more processes for risk management (ISO 14971)
▸ Determine quality objectives, verification, validation, monitoring, inspection and acceptance criteria
▸ Define records needed to show conformity
▸ Keep risk management records
ISO 14971 link
▸ Hazard identification
▸ Risk estimation & control
▸ Residual risk evaluation
▸ Production & post-production info

## Slide 31
7.2
Customer-Related Processes
SkillMed | ISO 13485:2016 Master Presentation Clause 7
31
1
Determine
Specified, implied, regulatory requirements; user training needs
2
Review
Before commitment — tenders, orders, contracts
3
Changes
Amend documents and inform personnel
4
Communicate
Product info, enquiries, feedback, complaints, advisory notices
5
Regulators
Communicate with authorities per regulations

## Slide 32
7.3.2 / 7.3.3
Design Planning & Inputs
SkillMed | ISO 13485:2016 Master Presentation Clause 7
32
D&D planning
▸ Stages and reviews at each stage
▸ Verification, validation & transfer activities
▸ Responsibilities & authorities
▸ Traceability of outputs to inputs
▸ Resources incl. competence
Design inputs
▸ Functional, performance, usability & safety
▸ Regulatory requirements & standards
▸ Risk management outputs
▸ Previous similar designs
▸ Complete, unambiguous, verifiable

## Slide 33
7.3.4 / 7.3.5
Design Outputs & Reviews
SkillMed | ISO 13485:2016 Master Presentation Clause 7
33
Outputs
▸ Meet input requirements
▸ Information for purchasing, production & service
▸ Product acceptance criteria
▸ Characteristics essential for safe use
Reviews
▸ Systematic at suitable stages
▸ Evaluate ability to meet requirements
▸ Include reps of functions concerned
▸ Independent specialist where needed
▸ Records of results & actions

## Slide 34
7.3.6 / 7.3.7
Verification vs Validation
SkillMed | ISO 13485:2016 Master Presentation Clause 7
34
Verification — built it right?
▸ Outputs meet inputs
▸ Tests, inspections, analyses
▸ Interfaces verified if connected to other devices
▸ Documented plans, methods, criteria
Validation — built the right device?
▸ Meets user needs & intended use
▸ Representative product (initial units/batches)
▸ Clinical evaluation / performance evaluation
▸ Completed before release for use

## Slide 35
7.3.8
Design Transfer
SkillMed | ISO 13485:2016 Master Presentation Clause 7
35
Process design
Manufacturing methods defined
Pilot build
Representative batches
Process validation
IQ / OQ / PQ
Docs & DMR
DMR, SOPs, inspections
Release
Production capability confirmed
Verify outputs are suitable for manufacturing before becoming final production specifications

## Slide 36
7.3.9
Control of Design Changes
SkillMed | ISO 13485:2016 Master Presentation Clause 7
36
▸ Identify and document design changes
▸ Determine significance of change to function, performance, usability, safety and regulatory requirements
▸ Review, verify, validate (as appropriate) and approve before implementation
▸ Evaluate effect on constituent parts, product in process or delivered, and risk management
▸ Assess regulatory notification need (e.g. Notified Body significant change)
Records
▸ Change request
▸ Impact assessment
▸ V&V evidence
▸ Approvals

## Slide 37
7.3.10
Design & Development Files
SkillMed | ISO 13485:2016 Master Presentation Clause 7
37
Record | Content | Purpose
Design & Development File (DHF) | Plans, inputs, outputs, reviews, V&V, changes | Shows design followed the plan
Device Master Record / MDF | Specifications, processes, labelling | Recipe to build the device
Device History Record / batch record | What was built, when, results | Evidence each lot meets DMR
Risk Management File | ISO 14971 analysis & controls | Safety rationale throughout life-cycle

## Slide 38
IEC 62304
Medical Device Software
SkillMed | ISO 13485:2016 Master Presentation Clause 7
38
1
Safety class
Class A, B or C based on harm
2
Development plan
Requirements, architecture, units, integration
3
Verification
Unit, integration and system testing
4
Maintenance
Problem resolution and change control
5
Cybersecurity
Threat modelling and SBOM where required
6
Usability
IEC 62366-1 user interface evaluation

## Slide 39
Audit pitfalls
Design Control Findings
SkillMed | ISO 13485:2016 Master Presentation Clause 7
39
1
Untraceable inputs
No trace matrix from user needs to tests
2
Late risk file
Risk analysis written after design freeze
3
Invalid samples
Validation on non-representative units
4
Weak reviews
No independent reviewer recorded
5
Uncontrolled changes
Changes without impact assessment
6
Incomplete DHF
Missing approvals or test reports

## Slide 40
05
Purchasing & Supplier Control
Clause 7.4 — evaluate, select, monitor and verify suppliers
40

## Slide 41
7.4.1
Purchasing Process
SkillMed | ISO 13485:2016 Master Presentation Clause 7.4
41
Criteria
Ability to supply, performance, risk
Evaluate
Questionnaire, audit, samples
Select & approve
Approved supplier list
Monitor
Re-evaluate performance
Act
Address deficiencies
Control is proportionate to the risk the supplied product poses to the finished device

## Slide 42
7.4.1
Supplier Categorisation
SkillMed | ISO 13485:2016 Master Presentation Clause 7.4
42
Category | Example | Typical controls
Critical | Sterilization, implant raw material, contract manufacturer | On-site audit, quality agreement, change notification
Major | Moulded parts, packaging, PCBs | Questionnaire, incoming inspection, CoC
Minor | Stationery, general consumables | Approved vendor only
Service | Calibration labs, test labs | Accreditation (ISO/IEC 17025) check

## Slide 43
7.4.2 / 7.4.3
Purchasing Info & Verification
SkillMed | ISO 13485:2016 Master Presentation Clause 7.4
43
Purchasing information
▸ Product specifications
▸ Acceptance & qualification requirements
▸ Personnel qualification
▸ QMS requirements
▸ Written agreement to notify changes
Verification of purchased product
▸ Inspection proportional to risk
▸ Act on discovered changes
▸ Verification at supplier when planned
▸ Records of verification

## Slide 44
Oversight
Outsourced Critical Services
SkillMed | ISO 13485:2016 Master Presentation Clause 7.4
44
1
Sterilization
Validated cycles (ISO 11135 / 11137), routine release records
2
Biocompatibility labs
ISO 10993 GLP testing, accredited labs
3
Contract manufacturers
Quality agreement, audits, DHR review
4
Testing labs
Scope of accreditation & raw data access

## Slide 45
06
Production & Service Provision
Clause 7.5 – 7.6 — controlled production, validation, traceability and calibration
45

## Slide 46
7.5.1
Control of Production
SkillMed | ISO 13485:2016 Master Presentation Clause 7.5
46
▸ Plan, carry out, monitor and control production to ensure conformity
▸ Documented procedures and methods, qualified infrastructure
▸ Monitoring and measurement of process parameters and product characteristics
▸ Labelling and packaging operations defined
▸ Record per batch: quantity manufactured and approved for distribution
Batch record
▸ Traceability
▸ Verified & approved
▸ Released by authorised person

## Slide 47
7.5.2 – 7.5.4
Cleanliness, Installation & Servicing
SkillMed | ISO 13485:2016 Master Presentation Clause 7.5
47
1
Cleanliness
Requirements when product is cleaned before sterilization or use
2
Installation
Acceptance criteria for installation and its verification
3
Servicing
Procedures, reference materials; analyse records for complaints
4
Sterile records
Records of sterilization process parameters per batch

## Slide 48
7.5.6
Process Validation
SkillMed | ISO 13485:2016 Master Presentation Clause 7.5
48
Plan
Protocol & criteria
IQ
Installed correctly
OQ
Works at limits
PQ
Consistent output
Revalidate
After changes / periodically
Validate any process whose output cannot be or is not verified by subsequent monitoring

## Slide 49
7.5.7
Sterile Medical Devices
SkillMed | ISO 13485:2016 Master Presentation Clause 7.5
49
▸ Document procedures for validation of sterilization and sterile barrier systems
▸ Validate before implementation and after product or process changes
▸ Common standards: ISO 11135 (EO), ISO 11137 (radiation), ISO 17665 (moist heat)
▸ Sterile barrier packaging per ISO 11607-1/-2
▸ Record results and conclusions of validation
Release
▸ Routine monitoring
▸ Biological / dosimetric indicators
▸ Parametric release where valid

## Slide 50
7.5.8 / 7.5.9
Identification & Traceability
SkillMed | ISO 13485:2016 Master Presentation Clause 7.5
50
Identification
▸ Identify product throughout realization
▸ Status relative to monitoring (accepted / rejected)
▸ Returned product distinguished
▸ UDI where required by regulation
Traceability
▸ Document procedures for traceability
▸ Implantables: traceability to components and environment
▸ Distributors keep distribution records
▸ Supports recall & field safety actions

## Slide 51
7.5.10 / 7.5.11
Customer Property & Preservation
SkillMed | ISO 13485:2016 Master Presentation Clause 7.5
51
1
Customer property
Identify, verify, protect; report loss or damage
2
Packaging
Design packaging to protect during processing and distribution
3
Storage
Special conditions controlled & recorded
4
Shelf-life
Limits and accelerated / real-time ageing evidence

## Slide 52
7.6
Monitoring & Measuring Equipment
SkillMed | ISO 13485:2016 Master Presentation Clause 7.5
52
▸ Calibrate or verify at intervals against traceable standards
▸ Adjust as necessary; identify calibration status
▸ Safeguard against adjustments that would invalidate results
▸ Assess validity of previous results when equipment is found out of tolerance
▸ Validate software used for monitoring and measurement
Calibration records
▸ Master list
▸ Certificates
▸ OOT assessments
▸ Labels

## Slide 53
07
Measurement, Analysis & Improvement
Clause 8 — feedback, complaints, audits, nonconformities, data and CAPA
53

## Slide 54
8.1 / 8.2.1
Feedback & Post-Market Surveillance
SkillMed | ISO 13485:2016 Master Presentation Clause 8
54
Feedback process
▸ Gather data from production and post-production
▸ Early warning of quality problems
▸ Input to risk management
▸ Input to product realization & improvement
PMS sources
▸ Complaints & service
▸ Field reports & literature
▸ Registries and vigilance databases
▸ Customer surveys

## Slide 55
8.2.2 / 8.2.3
Complaint Handling & Reporting
SkillMed | ISO 13485:2016 Master Presentation Clause 8
55
Receive
Record within timelines
Evaluate
Is it reportable?
Investigate
Or justify not
Act
Correction / CAPA
Report
Regulators & close
Reportable events go to authorities within regulatory timelines (e.g. EU MDR, FDA MDR, CDSCO)

## Slide 56
8.2.4
Internal Audits
SkillMed | ISO 13485:2016 Master Presentation Clause 8
56
▸ Audit at planned intervals to check conformity and effective implementation
▸ Programme based on status & importance of processes and previous results
▸ Auditors do not audit their own work
▸ Management takes corrections and corrective actions without undue delay
▸ Follow-up verifies actions taken and reports results
Audit cycle
▸ Plan
▸ Prepare checklist
▸ Conduct
▸ Report
▸ Follow up

## Slide 57
8.2.5 / 8.2.6 / 8.3
Monitoring & Nonconforming Product
SkillMed | ISO 13485:2016 Master Presentation Clause 8
57
Monitoring
▸ Process methods demonstrate planned results
▸ Product verified at stages per plans
▸ Release only after arrangements completed
▸ Record identity of releasing person
Nonconforming product
▸ Identify, document, segregate, evaluate
▸ Concession only with justification & regulatory check
▸ Action after delivery (advisory notice)
▸ Rework per documented procedure

## Slide 58
8.4
Analysis of Data — KPI Dashboard
SkillMed | ISO 13485:2016 Master Presentation Clause 8
58
Data source | Example KPI
Feedback | Complaint rate per 10,000 units
Product conformity | First-pass yield %
Process trends | Process capability (Cpk)
Suppliers | Supplier OTIF and rejection rate
Audits | Open findings past due
Service | Service reports per month

## Slide 59
8.5.2 / 8.5.3
CAPA
SkillMed | ISO 13485:2016 Master Presentation Clause 8
59
Identify
Nonconformity or risk
Investigate
5 Whys, fishbone, FTA
Plan action
Proportionate to effect
Verify
No adverse impact on compliance
Close
Effectiveness check & closure
Corrective = eliminate cause to prevent recurrence. Preventive = eliminate potential causes before occurrence.

## Slide 60
08
Certification & Roadmap
Audit process, common nonconformities and an implementation plan
60

## Slide 61
Stage 1 & 2
Certification Audit Process
SkillMed | ISO 13485:2016 Master Presentation Certification
61
Stage 1 — Readiness
▸ Documentation review
▸ Scope & site understanding
▸ Internal audit & management review done
▸ Readiness for stage 2
Stage 2 — Implementation
▸ On-site / remote evaluation
▸ Interviews & records sampling
▸ Nonconformity grading: major / minor
▸ Certificate on closure; annual surveillance; 3-year recertification

## Slide 62
Lessons
Top 10 Audit Nonconformities
SkillMed | ISO 13485:2016 Master Presentation Certification
62
# | Finding | # | Finding
1 | CAPA lacks root cause / effectiveness | 6 | Training effectiveness not evaluated
2 | Complaint not evaluated for reportability | 7 | Supplier controls not risk-based
3 | Process validation missing / outdated | 8 | Software not validated
4 | Design changes uncontrolled | 9 | Management review inputs missing
5 | Calibration OOT not assessed | 10 | Document & record control gaps

## Slide 63
Implementation Roadmap
SkillMed | ISO 13485:2016 Master Presentation Certification
63
1. Gap analysis
Current vs standard & regulations
2. Plan & train
Team, budget, awareness
3. Document
Manual, SOPs, MDF
4. Implement
Run processes, keep records
5. Audit & review
Internal audit, management review
6. Certify
Stage 1 & 2 with registrar
Typical timeline: 6 – 12 months depending on size, device class and maturity

## Slide 64
Key Takeaways
Regulatory first
Compliance & safety drive every clause
Risk everywhere
ISO 14971 links design, purchasing and production
Evidence
If it is not recorded, it did not happen
Learn more
Free ISO 13485 courses & certificates at skillmed.in
www.skillmed.in • info@skillmed.in • Educational material — verify against the official standard
64